The use of nanoparticles to improve the analytical capabilities of LIBS for solution sample analysis

V. Kiris1, A. Nevar1, N. Tarasenka1, N. Tarasenko1

M. Kuzmanović2, J. Petrović3, M. Momčilović3, J. Savović3

1Institute of Physics, National Academy of Sciences Minsk, Belarus

2University of Belgrade, Faculty of Physical Chemistry, Belgrade, Serbia

3Vinca Institute of Nuclear Sciences, Belgrade, Serbia

e-mail:lelas@vinca.rs

Laser-induced breakdown spectroscopy (LIBS) is a diagnostic technique based on optical emission spectroscopy, suitable for rapid, in situ, and multi-element analysis of all sample types. LIBS has demonstrated its greatest potential in the analysis of solids, both conducting and non-conducting. However, when it comes to quantitative LIBS analysis of liquids, there are still some drawbacks that limit its applicability, such as poor repeatability and low sensitivity. To some extent LIBS performance may be improved if directly performing LIBS on the solution is avoided by converting the liquid into a solid phase. Another recently proposed method to resolve the difficulties in LIBS analysis of liquid samples is the application of Nanoparticle - Enhanced LIBS (NELIBS) [1,2]. The aim of the present work was to examine the effect of copper oxide nanoparticles (NPs) on the analytical capabilities of LIBS analysis of lead in aqueous solutions.

The apparatus for NELIBS experiments comprises a laser source, a spectrograph (F=850 mm, grating 1800 groves/mm, F/11) coupled with a CCD, and the focusing and collecting optics, i.e. the same as used in conventional LIBS. A double pulse Q-switched Nd:YAG laser at 1064 nm, pulse duration 10 ns, pulse energy 50 mJ, interpulse delay 1 s was used. The spectra were registered in a time-integrated mode (integration time 1 ms). A glass plate was chosen as the substrate to ensure there is no interference between NPs and the substrate. The first type of samples was prepared by drying a droplet of a Pb solution (10 μl, concentration 100 ppm) on the glass substrate. The second type was prepared by drying a droplet of the copper oxide colloidal solution on the glass, prior to the deposition of the analyte solution.

Three sets of the LIBS spectra were acquired; spectra of glass with deposited CuO NPs, of glass with deposited sample solution of Pb, and of the glass with deposited NPs of CuO and the sample solution of Pb. The spectra were obtained in the 395-415 nm spectral range. Only in the third case, intense Pb lines appeared in the spectrum. Estimation of the limit of detection (LOD) for the resonance Pb line at 405.78 nm, without additional optimization of the measurement parameters, gave a value of 3.6 ppm. Regarding the fact that without the addition of the NPs layer no analyte signal was obtained, it can be concluded that the enhancement of the lead emission line intensity is caused by the presence of the NPs, i.e. by the lower plasma formation threshold and more efficient atomization of the analyte deposited on the NP layer.

The main benefits of using NELIBS for the analysis of solutions lie in its simplicity, speed, the small amount of solution needed (few microliters), and the ability to analyze very dense or slurry samples.

REFERENCES

[1] M. Dell'Aglio, R. Alrifai, A. De Giacomo, Spectrochim. Acta B 148, 105 (2018).

[2] A. De Giacomo, M. Dell’Aglio, R. Gaudiuso et al., J. Anal. At. Spectrom. 31, 1566 (2016).