Optical fiber curing of a dental composite:

a holographic, thermographic, and Raman study

E. Novta 1, T. Lainović 1, D. Grujić 2, S. Savić-Šević 2, E. Toth 3, Ž. Cvejić 3, L. Blažić 1,4, D. Pantelić 2

1University of Novi Sad, Faculty of Medicine, School of Dental Medicine, Novi Sad, Serbia

2University of Belgrade, Institute of Physics, Belgrade, Serbia

3University of Novi Sad, Faculty of Sciences, Department of Physics, Novi Sad, Serbia

4Dental Clinic of Vojvodina, Novi Sad, Serbia

e-mail: evgenije.novta@mf.uns.ac.rs

The main limitation of dental resin-based composites (RBC) is shrinkage during polymer chain formation. As a consequence, at the tooth-restoration interface polymerization shrinkage stress (PSS) is generated, impairing restoration longevity. The study aimed to introduce a novel photo-activation protocol using optical fibers inserted into the RBC, for reducing PSS and guiding light directly into deeper layers of the restoration.

Tooth models with standardized cavities were filled with the RBC (Filtek™ Bulk-fill Posterior, 3M ESPE), and the proposed experimental curing was designed as follows: in the first step, two optical fibers (ϕ 1.5 mm) connected with a commercial dental LED curing unit were inserted into the filling to cure the RBC from within; in the second step, fibers were extracted, remaining voids were filled with the RBC and final conventional curing was performed with the commercial light guide positioned above the sample. As a control, conventional curing was performed on a separate group of samples (n=15 models per curing protocol). Tooth model deformation as a secondary manifestation of PSS was measured in real-time using digital holographic interferometry. Simultaneous monitoring of RBC temperature change was conducted with an infrared thermal camera. The degree of monomer-to-polymer conversion (DC) was measured immediately after curing and after 24 h of dark storage, using Raman spectroscopy (n=5 models per curing protocol; three measurements points for each sample).

| Curing protocol | Radiant exposure (J/cm2) | Tooth model deformation (µm) | Temperature change (0C) | Time to reach the maximum temperature (s) | Immediate DC (%) | 24 h post-cure DC (%) |
| --- | --- | --- | --- | --- | --- | --- |
| Conventional | 20 | 13.4 | 10.7 | 7.5 | 41.4 | 46.3 |
| Experimental (1st + 2nd step) | 14.5 | 8.8 | 6.5 | 8.2 | 33.5 | 38.1 |

Table 1. Photo-activation parameters and results

The results showed a 34 % reduction of tooth model deformation using the experimental curing compared with the control. Such a finding was related to a lower temperature change (39 %), leading to a lower thermal shrinkage during the cooling process, as suggested by Palagummi et al. [1]. Furthermore, the time to reach the maximum exothermal temperature was prolonged compared with the control, allowing polymer chains to re-arrange and accommodate the volume reduction by viscous flow. Meanwhile, significantly lower immediate DC contributed to the reduction of model deformation as well. After 24 h of dark storage, the DC increased in both groups of samples as expected, but remained significantly lower for the ones cured using the experimental protocol compared with the control.

In conclusion, the proposed photo-activation using optical fibers was presented as beneficial for PSS reduction and could be easily implemented into dental clinical practice by coupling the fibers with standardized commercial light guides. However, it should be noted that measuring the irradiance of the selected light source is necessary to adequately determine exposure time since it was presented that insufficient radiant exposure (i.e. lower than suggested by the manufacturer) compromises the DC and therefore material biocompatibility and mechanics.

REFERENCES

[1] S.V. Palagummi, et al. Dent. Mater. 37(12), 2021: 1772-1782.